## Sheet2
Meter ID | Volunteer
TWI001 | TWI Staff
TWI002 | Felix Yu (Qian Liu) & Erick Yan
TWI003 | Ethan Wang (Qing Chang) & Frank Wang (Yunzhi Fan)
TWI004 | Ashritha & Bindu Nayanala
TWI005 | Todd Lanphear & Sharaya Gala
TWI006 | Dhruv & Rishabh Ramaswamy & Gabriel Pacheco
TWI007 | Warren Meade & Alexia Thanapalasingam & Nandana Kamath
TWI008 | Banu Mahalingam & Divya Seshadri
TWI009 | Daniel Gordon & Andy Weiss
TWI010 | Penny & Joyce Xia
TWI011 | Dan DiLollo
TWI012 | Bella Hu, Xin He & Jim Millonig
TWI013 | Allison Brown, Sue Wydick & Chris Ratanski
TWI014 | Plainsboro Preserve
TWI015 | Steve Johnson
TWI016 | Stephanie Liao (Kelly Liao) & Kristine Wang (Yilin Huang)
TWI017 | Anya Prasanna
TWI018 | Sophia Gonzales & Karolina Leleniewski
TWI019 | Annette Loveless & Noemi de la Puente
TWI020 | Patti Maslanka & Gloria Yu & Shresta Sathishkumar
TWI021 | Eric Dutaud & Laura Strong
TWI022 | Kari Brookhouse & Joseph Finnegan
TWI023 | Stolas Chong
TWI024 | Decommissioned
TWI025 | Unassigned

## Assignments
Meter ID | SN | Probe ID | SN | Volunteer | Type | Site1 | Site2 | Site3
TWI001 | 7260035101 | P001 |  | TWI Staff | Staff |  |  |  | AMB4 | 2 - Secondary | CRL1 | 0 - Retired/Inactive | MB5 | 0 - Retired/Inactive | SB4b | 1 - Priority
TWI002 | 7250069101 | P002 |  | Felix Yu (Qian Liu) & Erick Yan | Multi-Site | CL2 | CL4 | SB1 | AC1 | 1 - Priority | CTL1 | 3 - Tertiary | MDWL1 | 1 - Priority | SB5 | 3 - Tertiary
TWI003 | 7230005101 | P003 |  | Ethan Wang (Qing Chang) & Frank Wang (Yunzhi Fan) | Multi-Site | MRD | BB1 | CB3 | AC1a | 1 - Priority | CVL1 | 2 - Secondary | MI2 | 1 - Priority | SB6 | 2 - Secondary
TWI004 | 7250067101 | P004 |  | Ashritha & Bindu Nayanala | Multi-Site | BD2 | CR2 | RK2 | AC2 | 1 - Priority | DAM-6 | 1 - Priority | MI3 | 0 - Retired/Inactive | SB9 | 1 - Priority
TWI005 | 7260044101 | P005 |  | Todd Lanphear & Sharaya Gala | Multi-Site | CB4 | RB2 | BB2 | AC3 | 1 - Priority | DAM-7 | 1 - Priority | MI4 | 2 - Secondary | SBA | 1 - Priority
TWI006 | 7210028101 | P006 |  | Dhruv & Rishabh Ramaswamy & Gabriel Pacheco | Multi-Site | JC5 | FC1 | JC6 | AC4 | 1 - Priority | DB1 | 1 - Priority | MI4b | 1 - Priority | SBB | 1 - Priority
TWI007 | 7260005101 | P007 |  | Warren Meade & Alexia Thanapalasingam & Nandana Kamath | Multi-Site | AWL1 | SBA | SBC | AC5 | 1 - Priority | DB2 | 2 - Secondary | MO1 | 1 - Priority | SBC | 1 - Priority
TWI008 | 7250026101 | P008 |  | Banu Mahalingam & Divya Seshadri | Multi-Site | ROL1 | SB2 | 92 Elm Ridge | AC6 | 1 - Priority | DBA | 3 - Tertiary | MR1 | 1 - Priority | SBD | 1 - Priority
TWI009 | 7250012101 | P009 |  | Daniel Gordon & Andy Weiss | Multi-Site | MB3 | SB9 | HB3 | AC7 | 1 - Priority | DBB | 2 - Secondary | MR2 | 1 - Priority | SBE | 2 - Secondary
TWI010 | 7260037101 | P010 |  | Penny & Joyce Xia | Multi-Site | RK1 | BD4 | SYNL1 | AC8 | 1 - Priority | DE1 | 3- Tertiary | MR3 | 2 - Secondary | SC1 | 1 - Priority
TWI011 | 7210030101 | P011 |  | Dan DiLollo | Multi-Site | CC1 | GRPL1 | ATL1 | AC9 | 1 - Priority | DP3 | 1 - Priority | MR4 | 1 - Priority | SC2 | 1 - Priority
TWI012 | 7250065101 | P012 |  | Bella Hu, Xin He & Jim Millonig | Multi-Site | MR2 | DB2 | MR1 | APL1 | 1 - Priority | DS21-1 | 2 - Secondary | MR5 / DAM-3 | 1 - Priority | SC3 | 2 - Secondary
TWI013 | 7260048101 | P013 |  | Allison Brown, Sue Wydick & Chris Ratanski | Multi-Site | AX1 | SW1 | MO1 | ATL1 | 3 - Tertiary | EB1 | 1 - Priority | MR6 / DAM-4 | 1 - Priority | SC4 | 2 - Secondary
TWI014 | 7260010101 | P014 |  | Plainsboro Preserve | Multi-Site | PRP1 | DBA | DBB? | AWL1 | 2 - Secondary | ETL1 | 2 - Secondary | MR7 / DAM-9 | 1 - Priority | SC5 | 1 - Priority
TWI015 | 7260045101 | P015 |  | Steve Johnson | Multi-Site | HO3 | WP2 | WP4 | AX1 | 1 - Priority | FC1 | 1 - Priority | MR8 / DAM-10 | 1 - Priority | SC6 | 1 - Priority
TWI016 | 7260049101 | P016 |  | Stephanie Liao (Kelly Liao) & Kristine Wang (Yilin Huang) | Multi-Site | MR4 | CL5 | HA1 | AX2 | 0 - Retired/Inactive | GMP1 | 3 - Tertiary | MRA | 2 - Secondary | SHA | 1 - Priority
TWI017 | 7210007101 | P017 |  | Anya Prasanna | Multi-Site | RBA | BRL1 | CB5 | BB1 | 1 - Priority | GRPL1 | 2 - Secondary | MRB | 1 - Priority | SM1 | 1 - Priority
TWI018 | 7240015101 | P018 |  | Sophia Gonzales & Karolina Leleniewski | Multi-Site | SC6 | CVL1 | SYL1 | BB2 | 1 - Priority | HA1 | 1 - Priority | MRC | 1 - Priority | SM2 | 1 - Priority
TWI019 | 7250044101 | P019 |  | Annette Loveless & Noemi de la Puente | Multi-Site | LSC1 | COL1 | SC5 | BD1 | 0 - Retired/Inactive | HA2 | 1 - Priority | MRCL1 | 1 - Priority | SM3 | 2 - Secondary
TWI020 | 7240008101 | P020 |  | Patti Maslanka & Gloria Yu & Shresta Sathishkumar | Multi-Site | CR1 | PR1 | BK1 | BD1a | 1 - Priority | HB1 | 0 - Retired/Inactive | MRD | 1 - Priority | SPBP1 | 3 - Tertiary
TWI021 | 7240016101 | P021 |  | Eric Dutaud & Laura Strong | Multi-Site | MR3 | BD1a | PR2a | BD2 | 2 - Secondary | HB2 | 0 - Retired/Inactive | PB1 | 2 - Secondary | SPBP2 | 3 - Tertiary
TWI022 | 7210016101 | P022 |  | Kari Brookhouse & Joseph Finnegan | Multi-Site | AC1 | AC2 | AC3 | BD3 | 0 - Retired/Inactive | HB3 | 1 - Priority | PL1 | 1 - Priority | SW1 | 1 - Priority
TWI023 | 7230018101 | P023 |  | Stolas Chong | Multi-Site | APL1 | AC7 | AC6 | BD4 | 1 - Priority | HL1 | 0 - Retired/Inactive | PMRL1 | 3 - Tertiary | SYL1 | 2 - Secondary
TWI024 | 7240022101 | P024 |  | Decommissioned | Multi-Site |  |  |  | BD5 | 2 - Secondary | HL2 | 1 - Priority | PNVL1 | 3 - Tertiary | SYNL1 | 2 - Secondary
TWI025 | 7210029101 | P025 |  |  | Multi-Site |  |  |  | BD6 | 0 - Retired/Inactive | HL3 | 0 - Retired/Inactive | PP1 | 3 - Tertiary | TM1 | 1 - Priority
[9 empty] | BK1 | 2 - Secondary | HO1 | 1 - Priority | PR1 | 2 - Secondary | WB1 | 2 - Secondary
[9 empty] | BLDL1 | 3 - Tertiary | HO2 | 1 - Priority | PR2 | 0 - Retired/Inactive | WB2A | 2 - Secondary
LaMotte Users | [8 empty] | BRL1 | 2 - Secondary | HO3 | 1 - Priority | PR2a | 1 - Priority | WB2B | 0 - Retired/Inactive
Tom Miller |  | MI4b | [6 empty] | CAN-DP | 2 - Secondary | HO3a | 1 - Priority | PR3 | 2 - Secondary | WB2T | 0 - Retired/Inactive
Jim Gambino |  | SB2 | [6 empty] | CAN-LM | 2 - Secondary | JC0 | 2 - Secondary | PRP1 | 1 - Priority | WP1 | 2 - Secondary
Dave & Nicole Wagenblast |  | DB1 | [6 empty] | CB1 | 0 - Retired/Inactive | JC1 | 3 - Tertiary | PRP1a | 2 - Secondary | WP2 | 1 - Priority
Kirk & Jane Rothfuss & Mary Tulloss |  | EB1 | [6 empty] | CB2 | 0 - Retired/Inactive | JC2 | 3 - Tertiary | RB1 | 1 - Priority | WP3 | 2 - Secondary
Anthony & Glorianne Robbi |  | SB5 | [6 empty] | CB3 | 1 - Priority | JC3 | 2 - Secondary | RB2 | 1 - Priority | WP4 | 1 - Priority
Sonia, Maya & Franjo Ivancic |  | AC5 | [6 empty] | CB4 | 1 - Priority | JC4 | 3 - Tertiary | RBA | 1 - Priority
Ginger O'Shea & Sean Kirby |  | SC3 | [6 empty] | CB5 | 1 - Priority | JC4A | 0 - Retired/Inactive | RK1 | 1 - Priority
Winnie Spar |  | HA2 | [6 empty] | CC1 | 1 - Priority | JC4T | 0 - Retired/Inactive | RK2 | 1 - Priority
Doug Lischick |  | MRA | [6 empty] | CH1 | 0 - Retired/Inactive | JC5 | 1 - Priority | RO1 | 0 - Retired/Inactive
Andy Weiss |  | MB3 | [6 empty] | CL1 | 0 - Retired/Inactive | JC6 | 2 - Secondary | RO1a | 1 - Priority
Ted Chase |  | HB3 | [6 empty] | CL2 | 1 - Priority | JC9 | 0 - Retired/Inactive | RO2 | 1 - Priority
Pat Donahue |  | PL1 | [6 empty] | CL4 | 1 - Priority | LSC1 | 1 - Priority | ROL1 | 1 - Priority
Doug Kabay |  | HO1 | [6 empty] | CL5 | 1 - Priority | LSP1 | 0 - Retired/Inactive | SB1 | 1 - Priority
Kim Lavin |  | SC4 | [6 empty] | COL1 | 1 - Priority | MB1 | 0 - Retired/Inactive | SB2 | 1 - Priority
Rona Webster |  | BB1 | [6 empty] | CR1 | 1 - Priority | MB2 | 0 - Retired/Inactive | SB3 | 0 - Retired/Inactive
[9 empty] | CR2 | 1 - Priority | MB3 | 2 - Secondary | SB4 | 0 - Retired/Inactive
Note on U5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Formerly CL3
Note on I10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Formerly CL3
Note on B39: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Switched to meters
Note on Q40: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Formerly 128

## Sensors
Date Last Changed
Meter ID | SN | DO Sensor | pH Sensor | EC Sensor
TWI001 | 7260035101 | 2025-06-04 |  | 2025-10-25
TWI002 | 7250069101 |  |  | 2025-10-25
TWI003 | 7230005101 | 2024-01-01 |  | 2025-10-25
TWI004 | 7250067101 | 2024-04-01 | 2024-07-01 | 2025-10-25
TWI005 | 7260044101 |  | 2024-04-01 | 2025-10-25
TWI006 | 7210028101 |  |  | 2025-10-25
TWI007 | 7260005101 | 2025-10-25 | 2024-04-01 | 2025-10-25
TWI008 | 7250026101 |  | 2025-04-24 | 2025-10-25
TWI009 | 7250012101 |  |  | 2025-10-25
TWI010 | 7260037101 |  |  | 2025-10-25
TWI011 | 7210030101 |  |  | 2025-10-25
TWI012 | 7250065101 |  |  | 2025-10-25
TWI013 | 7260048101 |  | 2024-04-01 | 2025-10-25
TWI014 | 7260010101 |  |  | 2025-10-25
TWI015 | 7260045101 |  | 2025-10-25 | 2025-10-25
TWI016 | 7260049101 |  |  | 2025-10-25
TWI017 | 7210007101 |  | 2025-10-25 | 2025-10-25
TWI018 | 7240015101 |  | 2025-10-25 | 2025-10-25
TWI019 | 7250044101 |  |  | 2025-10-25
TWI020 | 7240008101 |  |  | 2025-10-25
TWI021 | 7240016101 |  | 2024-04-01 | 2025-10-25
TWI022 | 7210016101 |  | 2025-07-09 | 2025-10-25
TWI023 | 7230018101 |  | 2024-04-01 | 2025-10-25
TWI024 | 7240022101 | 2024-01-01
TWI025 | 7210029101 | 2025-07-09 | 2024-04-01 | 2025-10-25
Note on E17: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Sensor changed on 10/31 after issues calibrating at last sampling weekend

## Tracking
| Rate of Failure |  | 2024 |  |  |  | 2025 |  |  |  | 2026
|  |  | Jan | Apr | Jul | Oct | Jan | Apr | Jul | Oct | Jan | Apr | Jul | Oct
TWI001 | NA | T
| 37.5% | DO | C-F3 | C-F4 | P1 | F1 | P1 | P1 | P1 | P1
| 0% | pH | bar: 2024 Jan to 2025 Oct
| 50% | EC | F1 | P1 | P1 | P1 | C-F3 | F3 | P1 | C-F2
TWI002 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 12.5% | pH | F2 | P2 | P2 | P1 | P1 | P1 | P1 | P1
| 50% | EC | F1 | P1 | P1 | P1 | P1 | F3 | F4 | C-F2
TWI003 | NA | T
| 25% | DO | C-F3 | F4 | P2 | P1 | P1 | P1 | P1 | P1
| 0% | pH | bar: 2024 Jan to 2025 Oct
| 62.5% | EC | F1 | F1 | P2 | P1 | F2 | F3 | P3 | C-F2
TWI004 | NA | T
| 25% | DO | x | C-F4 | P3 | F1 | P1 | P3 | P1 | P1
| 0% | pH | x | P1 | C-P3 | P1 | P1 | P1 | P1 | P1
| 62.5% | EC | x | F1 | F1 | P1 | F2 | F1 | P4 | C-F2
TWI005 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 12.5% | pH | P1 | C-P2 | P1 | F2 | P1 | P1 | P1 | P1
| 62.5% | EC | F1 | F1 | P1 | P1 | C-F3 | F1 | P2 | C-F2
TWI006 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | bar: 2024 Jan to 2025 Oct
| 62.5% | EC | F1 | F1 | P1 | P1 | F2 | F1 | P4 | C-F2
TWI007 | NA | T
| 12.5% | DO | P1 | P1 | C-P4 | F1 | P1 | P2 | P3 | C-P5
| 0% | pH | P1 | C-P2 | P3 | P1 | P1 | P2 | P1 | P1
| 75% | EC | F1 | F1 | C-P2 | F1 | F2 | F3 | F4 | C-P5
TWI008 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | P1 | C-P2 | P1 | P1 | P1 | P1 | P1 | P3
| 75% | EC | F1 | F1 | P1 | P1 | F2 | F1 | C-F4 | C-F3
TWI009 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | bar: 2024 Jan to 2025 Oct
| 50% | EC | F1 | F1 | P1 | P1 | P1 | P1 | C-F4 | C-F2
TWI010 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 12.5% | pH | bar: 2024 Jan to 2025 Oct
| 37.5% | EC | F1 | P1 | P1 | P1 | P1 | F1 | P2 | C-F2
TWI011 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | bar: 2024 Jan to 2025 Oct
| 87.5% | EC | F1 | F1 | P1 | F1 | F2 | F3 | F4 | C-F2
TWI012 | NA | T
| 12.5% | DO | P1 | P1 | P2 | F1 | P1 | P1 | P1 | P2
| 12.5% | pH | P1 | P1 | P1 | F1 | P1 | P1 | P2 | P3
| 62.5% | EC | F1 | P1 | C-P2 | P1 | F2 | F3 | F4 | C-F3
TWI013 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | P1 | C-P2 | P1 | P1 | P1 | P1 | P1 | P1
| 75% | EC | F1 | F1 | C-P2 | P1 | F2 | F1 | F4 | C-F2
TWI014 | NA | T
| 12.5% | DO | P2 | F2 | P4 | P1 | P1 | P1 | P1 | P1
| 0% | pH | bar: 2024 Jan to 2025 Oct
| 75% | EC | F1 | F1 | P1 | F1 | F2 | P1 | C-F4 | C-F2
TWI015 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 12.5% | pH | P1 | P1 | P1 | F1 | C-P3 | P2 | P3 | P3
| 75% | EC | F1 | F1 | P1 | P1 | C-F3 | F1 | C-F4 | C-F3
TWI016 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | bar: 2024 Jan to 2025 Oct
| 75% | EC | F1 | x | P1 | F1 | F2 | F2 | C-F4 | C-F2
TWI017 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | P1 | P1 | P1 | P1 | P2 | P1 | P2 | C
| 50% | EC | F1 | F1 | P1 | P1 | F2 | F3 | P3 | C
TWI018 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 12.5% | pH | F2 | P2 | P1 | P1 | P1 | P3 | P2 | C
| 50% | EC | F1 | F1 | P1 | P1 | F2 | F3 | P4 | C
TWI019 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | bar: 2024 Jan to 2025 Oct
| 62.5% | EC | F1 | P1 | P1 | P1 | F2 | F2 | F4 | C-F2
TWI020 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 12.5% | pH | P1 | x | P1 | P1 | F2 | P1 | P1 | P1
| 75% | EC | F1 | x | P1 | F1 | C-F3 | F2 | C-F4 | C-F2
TWI021 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 0% | pH | P1 | C-P2 | P1 | P1 | P2 | P2 | P2 | P2
| 50% | EC | F1 | P1 | P1 | P1 | P1 | F2 | F4 | C-F2
TWI022 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 12.5% | pH | P1 | P1 | P1 | P1 | F2 | P1 | C-P4 | P1
| 75% | EC | F1 | F1 | P2 | P1 | F2 | F2 | F4 | C-F2
TWI023 | NA | T
| 0% | DO | bar: 2024 Jan to 2025 Oct
| 25% | pH | F2 | C-P2 | P1 | P1 | F2 | P1 | P1 | P1
| 25% | EC | P1 | P1 | P1 | P1 | P1 | F2 | P1 | C-F2
TWI024 | NA | T
| 25% | DO | C-F3 | F4 | P1 | P1 | P1 | P1 | P1 | x
| 12.5% | pH | F2 | P1 | P1 | P1 | P1 | P1 | P1 | x
| 37.5% | EC | F1 | P1 | P2 | P1 | F2 | F2 | P4 | x
TWI025 | NA | T
| 0% | DO | P1 | P1 | C-P4 | P1 | P1 | P1 | C-P3 | P1
| 0% | pH | P1 | C-P3 | P1 | P1 | P1 | P1 | P2 | P1
| 62.5% | EC | F1 | F1 | P1 | F1 | P1 | F1 | P2 | C-F2
Rate of Failure |  |  | 40% | 26.67% | 1.333% | 18.67% | 29.33% | 30.67% | 18.67% | 28% | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
DO
pH
EC
Note on J61: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Changed sensor on 10/31 after issues calibrating during last sampling event

## 2026 Testing
Meter Testing Procedures |  | December/January | [6 empty] | December/January | [6 empty] | December/January | [6 empty] | December/January
|  | Round 1 | [6 empty] | Round 2 | [6 empty] | Round 3 | [6 empty] | Round 4
| Troubleshooting | Test Date: | [6 empty] | Test Date: | [6 empty] | Test Date: | [6 empty] | Test Date:
| If a meter has failed one or more of the tests, the next step is to try to troubleshoot to correct the problem. In all cases, the first step is to turn off the meter, turn it back on and then recalibrate and retest. If the meter fails the second test, perform the following troubleshooting steps in the order in which they are written. Make sure to document the results of each round of calibration and testing. | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond.
|  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0
|  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0
|  | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0
|  | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0
|  | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0
|  | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0
|  | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0
|  | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0
|  | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0
|  | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0
|  | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0
| Troubleshoot DO | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0
| 1. Rehydrate the DO membrane, recalibrate & retest | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0
|  | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0
| 2. Reset to factory DO calibration, recalibrate & retest | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0
|  | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0
| 3. Replace the DO membrane and o-ring, recalibrate & retest | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0
|  | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0
| 4. Replace the DO sensor, install DO membrane and o-ring, reset to factory DO calibration, recalibrate & retest | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0
|  | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0
| Troubleshoot pH | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0
| 1. Rehydrate pH sensor in pH 4.01 buffer for at least 1 hour (or overnight), recalibrate & retest | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0
|  | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0
| 2. Reset to factory pH calibration, recalibrate & retest | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0
|  | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0
| 3. Replace the pH sensor, (rehydrate if necessary), reset to factory pH calibration, recalibrate & retest | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0
|  | NOTES: | [6 empty] | NOTES: | [6 empty] | NOTES: | [6 empty] | NOTES:
Troubleshoot Conductivity
1. Reset to factory conductivity calibration, recalibrate & retest
2. Replace conductivity sensor, reset to factory conductivity calibration, recalibrate & retest
|  | March/April | [6 empty] | March/April | [6 empty] | March/April | [6 empty] | March/April
|  | Round 1 | [6 empty] | Round 2 | [6 empty] | Round 3 | [6 empty] | Round 4
|  | Test Date: | [6 empty] | Test Date: | [6 empty] | Test Date: | [6 empty] | Test Date:
|  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond.
|  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0
|  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0
|  | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0
|  | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0
|  | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0
|  | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0
|  | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0
|  | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0
|  | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0
|  | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0
|  | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0
|  | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0
|  | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0
|  | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0
|  | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0
|  | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0
|  | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0
|  | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0
|  | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0
|  | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0
|  | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0
|  | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0
|  | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0
|  | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0
|  | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0
|  | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0
|  | NOTES: | [6 empty] | NOTES: | [6 empty] | NOTES: | [6 empty] | NOTES:
|  | June/July | [6 empty] | June/July | [6 empty] | June/July | [6 empty] | June/July
|  | Round 1 | [6 empty] | Round 2 | [6 empty] | Round 3 | [6 empty] | Round 4
|  | Test Date: | [6 empty] | Test Date: | [6 empty] | Test Date: | [6 empty] | Test Date:
|  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond.
|  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0
|  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0
|  | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0
|  | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0
|  | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0
|  | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0
|  | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0
|  | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0
|  | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0
|  | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0
|  | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0
|  | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0
|  | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0
|  | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0
|  | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0
|  | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0
|  | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0
|  | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0
|  | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0
|  | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0
|  | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0
|  | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0
|  | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0
|  | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0
|  | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0
|  | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0
|  | NOTES: | [6 empty] | NOTES: | [6 empty] | NOTES: | [6 empty] | NOTES:
|  | September/October | [6 empty] | September/October | [6 empty] | September/October | [6 empty] | September/October
|  | Round 1 | [6 empty] | Round 2 | [6 empty] | Round 3 | [6 empty] | Round 4
|  | Test Date: | [6 empty] | Test Date: | [6 empty] | Test Date: | [6 empty] | Test Date:
|  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond.
|  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0 |  |  |  | 0 | 0 |  | 0
|  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0 | TWI001 |  | 0 |  |  |  | 0
|  | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0 | TWI002 |  | 0 |  |  |  | 0
|  | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0 | TWI003 |  | 0 |  |  |  | 0
|  | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0 | TWI004 |  | 0 |  |  |  | 0
|  | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0 | TWI005 |  | 0 |  |  |  | 0
|  | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0 | TWI006 |  | 0 |  |  |  | 0
|  | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0 | TWI007 |  | 0 |  |  |  | 0
|  | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0 | TWI008 |  | 0 |  |  |  | 0
|  | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0 | TWI009 |  | 0 |  |  |  | 0
|  | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0 | TWI010 |  | 0 |  |  |  | 0
|  | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0 | TWI011 |  | 0 |  |  |  | 0
|  | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0 | TWI012 |  | 0 |  |  |  | 0
|  | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0 | TWI013 |  | 0 |  |  |  | 0
|  | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0 | TWI014 |  | 0 |  |  |  | 0
|  | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0 | TWI015 |  | 0 |  |  |  | 0
|  | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0 | TWI016 |  | 0 |  |  |  | 0
|  | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0 | TWI017 |  | 0 |  |  |  | 0
|  | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0 | TWI018 |  | 0 |  |  |  | 0
|  | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0 | TWI019 |  | 0 |  |  |  | 0
|  | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0 | TWI020 |  | 0 |  |  |  | 0
|  | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0 | TWI021 |  | 0 |  |  |  | 0
|  | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0 | TWI022 |  | 0 |  |  |  | 0
|  | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0 | TWI023 |  | 0 |  |  |  | 0
|  | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0 | TWI024 |  | 0 |  |  |  | 0
|  | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0 | TWI025 |  | 0 |  |  |  | 0
|  | NOTES: | [6 empty] | NOTES: | [6 empty] | NOTES: | [6 empty] | NOTES:

## 2025 Testing
Meter Testing Procedures |  | JANUARY | [9 empty] | JANUARY | [9 empty] | JANUARY
|  | Round 1 | [9 empty] | Round 2 | [9 empty] | Round 3
| Troubleshooting | Test Date: | 2025-01-09 | [8 empty] | Test Date: | 2025-01-09 | [8 empty] | Test Date: | 2025-01-09
| If a meter has failed one or more of the tests, the next step is to try to troubleshoot to correct the problem. In all cases, the first step is to turn off the meter, turn it back on and then recalibrate and retest. If the meter fails the second test, perform the following troubleshooting steps in the order in which they are written. Make sure to document the results of each round of calibration and testing. | Standards: | pH | DO% | DO ppm | Cond. 1 | ±1% cond .1 | Cond. 2 | ± 1% cond. 2 |  |  | Standards: | pH | DO% | DO ppm | Cond. 1 | ±1% cond .1 | Cond. 2 | ± 1% cond. 2 |  |  | Standards: | pH | DO% | DO ppm | Cond. 1 | ±1% cond .1 | Cond. 2 | ± 1% cond. 2
|  |  | 7 | 0 | 0 | 720 | 7.2 |  | 0 |  |  |  | 7 | 0 | 0 | 720 | 7.2 | [5 empty] | 7 | 0 | 0 | 720 | 7.2 |  | 0
|  | Meter | pH | pH diff. | DO% | DO ppm | Cond. 1 | Cond. 1 Diff | Cond. 2 | Cond. 2 Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. 1 | Cond. 1 Diff | Cond. 2 | Cond. 2 Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. 1 | Cond. 1 Diff | Cond. 2 | Cond. 2 Diff.
|  | x | 6.99 | -0.01 | 2.2% | 0.18 | 626 | -94 |  | 0 |  | TWI001* |  |  |  |  | 357 | -363 |  | 0 |  | TWI001 |  |  |  |  | 265 | -455 |  | 0
|  | TWI002 | 7.17 | 0.17 | 0% | 0 | 715 | -5 |  | 0 |  | TWI002 | [7 empty] | 0 |  | TWI002 | [7 empty] | 0
|  | TWI003 | 7.01 | 0.01 | 0.6% | 0.05 | 828 | 108 |  | 0 |  | TWI003 |  |  |  |  | 341 | -379 |  | 0 |  | TWI003 | [7 empty] | 0
|  | TWI004 | 7.13 | 0.13 | 0.8% | 0.07 | 444 | -276 |  | 0 |  | TWI004 |  |  |  |  | 375 | -345 |  | 0 |  | TWI004 | [7 empty] | 0
|  | TWI005 | 6.93 | -0.07 | 0% | 0 | 434 | -286 |  | 0 |  | TWI005* |  |  |  |  | 308 | -412 |  | 0 |  | TWI005 |  |  |  |  | 264 | -456 |  | 0
|  | TWI006 | 7.11 | 0.11 | 0% | 0 | 777 | 57 |  | 0 |  | TWI006 |  |  |  |  | 340 | -380 |  | 0 |  | TWI006 | [7 empty] | 0
|  | TWI007 | 7.1 | 0.1 | 2% | 0.17 | 486 | -234 |  | 0 |  | TWI007 |  |  |  |  | 294 | -426 |  | 0 |  | TWI007 | [7 empty] | 0
|  | TWI008 | 7.15 | 0.15 | 0% | 0 | 737 | 17 |  | 0 |  | TWI008 |  |  |  |  | 323 | -397 |  | 0 |  | TWI008 | [7 empty] | 0
|  | TWI009 | 7.13 | 0.13 | 0% | 0 | 726 | 6 |  | 0 |  | TWI009 | [7 empty] | 0 |  | TWI009 | [7 empty] | 0
|  | TWI010 | 7.1 | 0.1 | 0% | 0 | 719 | -1 |  | 0 |  | TWI010 | [7 empty] | 0 |  | TWI010 | [7 empty] | 0
| Troubleshoot DO | TWI011 | 7.14 | 0.14 | 0% | 0 | 618 | -102 |  | 0 |  | TWI011 |  |  |  |  | 353 | -367 |  | 0 |  | TWI011 | [7 empty] | 0
| 1. Rehydrate the DO membrane, recalibrate & retest | TWI012 | 7.05 | 0.05 | 2.1% | 0.18 | 627 | -93 |  | 0 |  | TWI012 |  |  |  |  | 358 | -362 |  | 0 |  | TWI012 | [7 empty] | 0
|  | TWI013 | 7.14 | 0.14 | 0% | 0 | 616 | -104 |  | 0 |  | TWI013 |  |  |  |  | 357 | -363 |  | 0 |  | TWI013 | [7 empty] | 0
| 2. Reset to factory DO calibration, recalibrate & retest | TWI014 | 7 | 0 | 0% | 0 | 408 | -312 |  | 0 |  | TWI014 |  |  |  |  | 354 | -366 |  | 0 |  | TWI014 | [7 empty] | 0
|  | TWI015 | 7.3 | 0.3 | 0% | 0 | 767 | 47 |  | 0 |  | TWI015** | 7.25 | 0.25 |  |  | 315 | -405 |  | 0 |  | TWI015 | 7.17 | 0.17 |  |  | 268 | -452 |  | 0
| 3. Replace the DO membrane and o-ring, recalibrate & retest | TWI016 | 7.15 | 0.15 | 0% | 0 | 775 | 55 |  | 0 |  | TWI016 |  |  |  |  | 347 | -373 |  | 0 |  | TWI016 | [7 empty] | 0
|  | TWI017 | 7.26 | 0.26 | 0% | 0 | 770 | 50 |  | 0 |  | TWI017 | 7.17 | 0.17 |  |  | 325 | -395 |  | 0 |  | TWI017 | [7 empty] | 0
| 4. Replace the DO sensor, install DO membrane and o-ring, reset to factory DO calibration, recalibrate & retest | TWI018 | 7.1 | 0.1 | 0% | 0 | 789 | 69 |  | 0 |  | TWI018 |  |  |  |  | 337 | -383 |  | 0 |  | TWI018 | [7 empty] | 0
|  | TWI019 | 7.16 | 0.16 | 0% | 0 | 818 | 98 |  | 0 |  | TWI019 |  |  |  |  | 330 | -390 |  | 0 |  | TWI019 | [7 empty] | 0
| Troubleshoot pH | TWI020 | 7.24 | 0.24 | 0% | 0 | 518 | -202 |  | 0 |  | TWI020* | 7.24 | 0.24 |  |  | 321 | -399 |  | 0 |  | TWI020 |  |  |  |  | 289 | -431 |  | 0
| 1. Rehydrate pH sensor in pH 4.01 buffer for at least 1 hour (or overnight), recalibrate & retest | TWI021 | 7.26 | 0.26 | 0% | 0 | 724 | 4 |  | 0 |  | TWI021 | 7.12 | 0.12 | [5 empty] | 0 |  | TWI021 | [7 empty] | 0
|  | TWI022 | 7.34 | 0.34 | 0% | 0 | 840 | 120 |  | 0 |  | TWI022 | 7.26 | 0.26 |  |  | 323 | -397 |  | 0 |  | TWI022 | [7 empty] | 0
| 2. Reset to factory pH calibration, recalibrate & retest | TWI023 | 7.33 | 0.33 | 0% | 0 | 721 | 1 |  | 0 |  | TWI023 | 7.25 | 0.25 | [5 empty] | 0 |  | TWI023 | [7 empty] | 0
|  | TWI024 | 6.95 | -0.05 | 0% | 0 | 687 | -33 |  | 0 |  | TWI024 |  |  |  |  | 331 | -389 |  | 0 |  | TWI024 | [7 empty] | 0
| 3. Replace the pH sensor, (rehydrate if necessary), reset to factory pH calibration, recalibrate & retest | TWI025 | 7.13 | 0.13 | 0% | 0 | 713 | -7 |  | 0 |  | TWI025 | [7 empty] | 0 |  | TWI025 | [7 empty] | 0
|  | NOTES: | Replaced DO caps. All meters calibrated at 100% DO, $.01 & 10 pH & 100 µS/cm then tested at 7 pH, 0% DO & 74 & 447 µS/cm. All meters passed DO test in 1st round. TWI015, TWI017, TWI020, TWI021, TWI022 & TWI023 failed pH. All but TWI002, TWO009, TWI010 & TWI025 failed conductivity. | [8 empty] | NOTES: | TWI015, TWI017, TWI020, TWI021, TWI022, & TWI023 were recalibrated at 4.01 & 10 pH and then restested at 7 pH. Only TWI017 & TWI021 passed pH test in the second round. All meters that failed conductivity in first round were recalibrated at 100 µS/cm and then restested at 720 µS/cm. None passed second round conductivity tests. | [8 empty] | NOTES: | Limited sensors on hand for replacement. Changed EC sensors on TWI001, TWI005, TWI015, & TWI020, reset to factory calibration and then recalibrated at 100 µS/cm. They all failed the 3rd round conductivity test at 720 µS/cm. Changed pH sensor on TWI015, reset to factory calibration and then recalibrated at 4.01 and 10 pH. It passed the 3rd round pH test at 7 pH.
Troubleshoot Conductivity
1. Reset to factory conductivity calibration, recalibrate & retest
2. Replace conductivity sensor, reset to factory conductivity calibration, recalibrate & retest
|  | March/April | [7 empty] | March/April | [7 empty] | March
|  | Round 1 | [7 empty] | Round 2 | [7 empty] | Round 3
|  | Test Date: | 2025-04-09 | [6 empty] | Test Date: | [7 empty] | Test Date:
|  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond.
|  |  | 7.01 | 0 | 0 | 718 | 7.18 |  |  |  | 7.01 | 0 | 0 | 718 | 7.18 |  |  |  | 7.01 | 0 | 0 | 718 | 7.18
|  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  | TWI001 | 7.1 | 0.1 | 3% | 0.26 | 678 | -40 |  | TWI001 | [7 empty] | TWI001
|  | TWI002 | 10.08 | 3.08 | 0% | 0 | 593 | -125 |  | TWI002 | 6.95 | -0.06 |  |  | 567 | -151 |  | TWI002 |  |  |  |  | 515 | -203
|  | TWI003 | 7.12 | 0.12 | 2.3% | 0.2 | 646 | -72 |  | TWI003 |  |  |  |  | 544 | -174 |  | TWI003 |  |  |  |  | 520 | -198
|  | TWI004 | 7.05 | 0.05 | 10.3% | 0.87 | 663 | -55 |  | TWI004 |  |  | 6.9% | 0.79 |  |  |  | TWI004 |  |  | 2.9% | 0.25
|  | TWI005 | 6.84 | -0.16 | 0% | 0 | 670 | -48 |  | TWI005 | [7 empty] | TWI005
|  | TWI006 | 7.18 | 0.18 | 0% | 0 | 655 | -63 |  | TWI006 | [7 empty] | TWI006
|  | TWI007 | 7.28 | 0.28 | 6.5% | 0.57 | 645 | -73 |  | TWI007 | 7.17 | 0.16 | 2% | 0.23 | 599 | -119 |  | TWI007 |  |  |  |  | 565 | -153
|  | TWI008 | 7.2 | 0.2 | 0% | 0 | 620 | -98 |  | TWI008 | [7 empty] | TWI008
|  | TWI009 | 7.19 | 0.19 | 0% | 0 | 713 | -5 |  | TWI009 | [7 empty] | TWI009
|  | TWI010 | 7.18 | 0.18 | 0% | 0 | 681 | -37 |  | TWI010 | [7 empty] | TWI010
|  | TWI011 | 7.11 | 0.11 | 0% | 0 | 655 | -63 |  | TWI011 |  |  |  |  | 601 | -117 |  | TWI011 |  |  |  |  | 540 | -178
|  | TWI012 | 7.2 | 0.2 | 1.7% | 0.15 | 532 | -186 |  | TWI012 |  |  |  |  | 598 | -120 |  | TWI012 |  |  |  |  | 492 | -226
|  | TWI013 | 7.2 | 0.2 | 0% | 0 | 682 | -36 |  | TWI013 | [7 empty] | TWI013
|  | TWI014 | 7 | 0 | 2.6% | 0.22 | 717 | -1 |  | TWI014 | [7 empty] | TWI014
|  | TWI015* | 6.64 | -0.36 | 0% | 0 | 697 | -21 |  | TWI015 | 7.2 | 0.19 | [5 empty] | TWI015
|  | TWI016 | 7.18 | 0.18 | 0% | 0 | 580 | -138 |  | TWI016 |  |  |  |  | 620 | -98 |  | TWI016 |  |  |  |  | 540 | -178
|  | TWI017 | 7.03 | 0.03 | 0% | 0 | 578 | -140 |  | TWI017 |  |  |  |  | 585 | -133 |  | TWI017 |  |  |  |  | 507 | -211
|  | TWI018 | 7.6 | 0.6 | 0% | 0 | 600 | -118 |  | TWI018 | 7.78 | 0.77 |  |  | 512 | -206 |  | TWI018 | 7.13 | 0.13 |  |  | 476 | -242
|  | TWI019 | 7.15 | 0.15 | 0% | 0 | 585 | -133 |  | TWI019 |  |  |  |  | 696 | -22 |  | TWI019
|  | TWI020 | 7.01 | 0.01 | 0% | 0 | 535 | -183 |  | TWI020 |  |  |  |  | 782 | 64 |  | TWI020
|  | TWI021 | 7.26 | 0.26 | 0% | 0 | 580 | -138 |  | TWI021 | 7.15 | 0.14 |  |  | 688 | -30 |  | TWI021
|  | TWI022 | 7.05 | 0.05 | 3.1% | 0.27 | 512 | -206 |  | TWI022 |  |  |  |  | 652 | -66 |  | TWI022
|  | TWI023 | 7.1 | 0.1 | 0% | 0 | 616 | -102 |  | TWI023 |  |  |  |  | 675 | -43 |  | TWI023
|  | TWI024 | 7 | 0 | 0% | 0 | 473 | -245 |  | TWI024 |  |  |  |  | 690 | -28 |  | TWI024
|  | TWI025 | 6.97 | -0.03 | 0% | 0 | 666 | -52 |  | TWI025 | [7 empty] | TWI025
|  | NOTES: | Replaced DO caps. All meters calibrated at 100% DO, 4.01 & 10 pH & 100 µS/cm then tested at 7 pH, 0% DO & 718 µS/cm. *Bold & red conductivity is greater than 10% off.* | [6 empty] | NOTES: | [7 empty] | NOTES:
|  | June/July | [7 empty] | June/July | [7 empty] | June/July | [7 empty] | June/July
|  | Round 1 | [7 empty] | Round 2 | [7 empty] | Round 3 | [7 empty] | Round 4
|  | Test Date: | 2025-07-08 | [6 empty] | Test Date: | 2025-07-09 | [6 empty] | Test Date: | 2025-07-09 | [6 empty] | Test Date: | 2025-07-10
|  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond.
|  |  | 7.01 | 0 | 0 | 718 | 7.18 |  |  |  | 7.01 | 0 | 0 | 718 | 7.18 |  |  |  | 7.01 | 0 | 0 | 718 | 7.18 |  |  |  | 7.01 | 0 | 0 | 718 | 7.18
|  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  | TWI001 | 6.98 | -0.02 | 2.9% | 0.24 | 725 | 7 |  | TWI001 | [7 empty] | TWI001 | [7 empty] | TWI001
|  | TWI002 | 7.12 | 0.12 | 0% | 0 | 526 | -192 |  | TWI002 |  |  |  |  | 641 | -77 |  | TWI002 |  |  |  |  | 630 | -88 |  | TWI002 |  |  |  |  | 665 | -53
|  | TWI003 | 7.07 | 0.07 | 3.6% | 0.3 | 748 | 30 |  | TWI003 |  |  |  |  | 730 | 12 |  | TWI003 |  |  |  |  | 724 | 6 |  | TWI003
|  | TWI004 | 7.2 | 0.2 | 1.1% | 0.09 | 521 | -197 |  | TWI004 |  |  |  |  | 614 | -104 |  | TWI004 |  |  |  |  | 622 | -96 |  | TWI004 |  |  |  |  | 719 | 1
|  | TWI005 | 6.9 | -0.1 | 0% | 0 | 532 | -186 |  | TWI005 |  |  |  |  | 719 | 1 |  | TWI005 | [7 empty] | TWI005
|  | TWI006 | 6.85 | -0.15 | 0% | 0 | 584 | -134 |  | TWI006 |  |  |  |  | 653 | -65 |  | TWI006 |  |  |  |  | 645 | -73 |  | TWI006 |  |  |  |  | 720 | 2
|  | TWI007 | 7.2 | 0.2 | 5.1% | 0.43 | 635 | -83 |  | TWI007 |  |  | 3.7% | 0.31 | 579 | -139 |  | TWI007 |  |  | 2.9% | 0.24 | 736 | 18 |  | TWI007 |  |  |  |  | 449 | -269
|  | TWI008 | 7.11 | 0.11 | 0% | 0 | 639 | -79 |  | TWI008 |  |  |  |  | 640 | -78 |  | TWI008 |  |  |  |  | 671 | -47 |  | TWI008 |  |  |  |  | 438 | -280
|  | TWI009 | 7.13 | 0.13 | 2.2% | 0.18 | 779 | 61 |  | TWI009 |  |  |  |  | 750 | 32 |  | TWI009 |  |  |  |  | 591 | -127 |  | TWI009 |  |  |  |  | 480 | -238
|  | TWI010 | 7.05 | 0.05 | 0% | 0 | 726 | 8 |  | TWI010 |  |  |  |  | 722 | 4 |  | TWI010 | [7 empty] | TWI010
|  | TWI011 | 7.11 | 0.11 | 0% | 0 | 645 | -73 |  | TWI011 |  |  |  |  | 731 | 13 |  | TWI011 |  |  |  |  | 709 | -9 |  | TWI011 |  |  |  |  | 495 | -223
|  | TWI012 | 7.23 | 0.23 | 1.1% | 0.09 | 602 | -116 |  | TWI012 | 7.15 | 0.14 |  |  | 644 | -74 |  | TWI012 |  |  |  |  | 647 | -71 |  | TWI012 |  |  |  |  | 653 | -65
|  | TWI013 | 7.13 | 0.13 | 0% | 0 | 682 | -36 |  | TWI013 |  |  |  |  | 678 | -40 |  | TWI013 |  |  |  |  | 655 | -63 |  | TWI013 |  |  |  |  | 678 | -40
|  | TWI014 | 7.1 | 0.1 | 0% | 0 | 581 | -137 |  | TWI014 |  |  |  |  | 543 | -175 |  | TWI014 |  |  |  |  | 615 | -103 |  | TWI014 |  |  |  |  | 637 | -81
|  | TWI015 | 7.5 | 0.5 | 0% | 0 | 500 | -218 |  | TWI015 | 7.22 | 0.21 |  |  | 601 | -117 |  | TWI015 | 7.18 | 0.18 |  |  | 577 | -141 |  | TWI015 |  |  |  |  | 691 | -27
|  | TWI016 | 7.2 | 0.2 | 0% | 0 | 200-blinking | #VALUE! |  | TWI016 |  |  |  |  | 546 | -172 |  | TWI016 |  |  |  |  | 563 | -155 |  | TWI016 |  |  |  |  | 486 | -232
|  | TWI017 | 7.27 | 0.27 | 0% | 0 | 518 | -200 |  | TWI017 | 7.2 | 0.19 |  |  | 582 | -136 |  | TWI017 |  |  |  |  | 725 | 7 |  | TWI017
|  | TWI018 | 8.28 | 1.28 | 0% | 0 | 200-blinking | #VALUE! |  | TWI018 | 7.01 | 0 |  |  | 642 | -76 |  | TWI018 |  |  |  |  | 638 | -80 |  | TWI018 |  |  |  |  | 722 | 4
|  | TWI019 | 7.17 | 0.17 | 0% | 0 | 461 | -257 |  | TWI019 |  |  |  |  | 629 | -89 |  | TWI019 |  |  |  |  | 587 | -131 |  | TWI019 |  |  |  |  | 647 | -71
|  | TWI020 | 7.19 | 0.19 | 0% | 0 | 645 | -73 |  | TWI020 |  |  |  |  | 534 | -184 |  | TWI020 |  |  |  |  | 522 | -196 |  | TWI020 |  |  |  |  | 511 | -207
|  | TWI021 | 7.43 | 0.43 | 2% | 0.17 | 651 | -67 |  | TWI021 | 7.17 | 0.16 |  |  | 514 | -204 |  | TWI021 |  |  |  |  | 772 | 54 |  | TWI021 |  |  |  |  | 699 | -19
|  | TWI022 | 7.22 | 0.22 | 0% | 0 | 658 | -60 |  | TWI022 | 7.33 | 0.32 |  |  | 651 | -67 |  | TWI022 | 7.31 | 0.31 |  |  | 681 | -37 |  | TWI022 | 7.16 | 0.16 |  |  | 454 | -264
|  | TWI023 | 7.16 | 0.16 | 0% | 0 | 711 | -7 |  | TWI023 | [7 empty] | TWI023 | [7 empty] | TWI023
|  | TWI024 | 7.01 | 0.01 | 0% | 0 | 200-blinking | #VALUE! |  | TWI024 |  |  |  |  | 640 | -78 |  | TWI024 |  |  |  |  | 670 | -48 |  | TWI024 |  |  |  |  | 713 | -5
|  | TWI025 | 7.24 | 0.24 | climbing | climbing | 685 | -33 |  | TWI025 | 7.18 | 0.17 | climbing | climbing | 723 | 5 |  | TWI025 |  |  | 1.6% | 0.13 |  |  |  | TWI025
|  | NOTES: | Replaced DO caps. All meters calibrated at 100% DO, 4.01 & 10 pH & 100 µS/cm then tested at 7 pH, 0% DO & 718 µS/cm. All but TWI001 and TWI023 failed conductivity. TWI007 and TWI025 failed DO. TWI012, TWI015, TWI017, TWI018, TWI021, TWI022 and TWI025 failed pH. TWI018 and TWI006 showed "check sensor" errors on first pH calibration. | [6 empty] | NOTES: | Working with three interns, each calibrating and testing for one of the parameters. Meters that failed were recalibrated & retested. TWI018 showed "check sensor" and "contaminated buffer" error on pH calibration but passed second test and allowed pH 7.01 calibration. | [6 empty] | NOTES: | Meters that failed the second test were reset to factory calibration, recalibrated & restested. Replaced DO sensor on TWI025, reset to factory calibration but would not recalibrate. Replaced pH sensor on TWI022, left to soak, then reset to factory calibration, recalibrated & restested. | [6 empty] | NOTES: | Meters were reset to factory conductivity calibration, recalibrated and retested. All but TWI004, TWI006, TWI018 and TWI024 failed again. Changed EC sensors on TWI009, TWI014, TWI015, TWI016, TWI020, and TWI008.TWI025 passed the DO test after resetting and recalibrating.
|  | September/October | [7 empty] | September/October | [7 empty] | Semptember/October | [7 empty] | September/October
|  | Round 1 | [7 empty] | Round 2 | [7 empty] | Round 3 | [7 empty] | Round 4
|  | Test Date: | 2025-10-09 | [6 empty] | Test Date: | 10/9/2025 & 10/10/25 | [6 empty] | Test Date: | 2025-10-10 | [6 empty] | Test Date: | 10/12/2025 & 10/14/25
|  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ±1% cond.
|  |  | 7.01 | 0 | 0 | 718 | 7.18 |  |  |  | 7.01 | 0 | 0 | 718 | 7.18 |  |  |  | 7.01 | 0 | 0 | 718 | 7.18 |  |  |  | 7.01 | 0 | 0 | 718 | 7.18
|  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  | TWI001 | 6.97 | -0.03 | 1.2% | 0.1 | 708 | -10 |  | TWI001 |  |  |  |  | 777 | 59 |  | TWI001 | [7 empty] | TWI001
|  | TWI002 | 7.07 | 0.07 | 0% | 0 | blinking | #VALUE! |  | TWI002 |  |  |  |  | 530 | -188 |  | TWI002 | [7 empty] | TWI002
|  | TWI003 | 7.17 | 0.17 | 2.7% | 0.23 | 695 | -23 |  | TWI003 |  |  |  |  | 1229 | 511 |  | TWI003 | [7 empty] | TWI003
|  | TWI004 | 7.1 | 0.1 | 0.09% | 0.08 | 645 | -73 |  | TWI004 |  |  |  |  | 819 | 101 |  | TWI004 | [7 empty] | TWI004
|  | TWI005 | 6.94 | -0.06 | 0% | 0 | blinking | #VALUE! |  | TWI005 |  |  |  |  | 651 | -67 |  | TWI005 | [7 empty] | TWI005
|  | TWI006 | 7.15 | 0.15 | 0% | 0 | 564 | -154 |  | TWI006 |  |  |  |  | 769 | 51 |  | TWI006 | [7 empty] | TWI006
|  | TWI007 | 7.15 | 0.15 | 5.5% | 0.48 | 695 | -23 |  | TWI007 |  |  | 5.2% | 0.45 | 607 | -111 |  | TWI007 |  |  | 10.8% | 0.97 | 555 | -163 |  | TWI007 |  |  | 1.4% | 0.12 | 718 | 0
|  | TWI008 | 7.32 | 0.32 | 0% | 0 | 580 | -138 |  | TWI008 | 7.36 | 0.35 |  |  | 646 | -72 |  | TWI008 | 7.1 | 0.1 |  |  | 560 | -158 |  | TWI008
|  | TWI009 | 7.14 | 0.14 | 1.6% | 0.14 | 513 | -205 |  | TWI009 |  |  |  |  | 686 | -32 |  | TWI009 | [7 empty] | TWI009
|  | TWI010 | 7.2 | 0.2 | 0% | 0 | 732 | 14 |  | TWI010 |  |  |  |  | 412 | -306 |  | TWI010 | [7 empty] | TWI010
|  | TWI011 | 7.11 | 0.11 | 0% | 0 | 673 | -45 |  | TWI011 |  |  |  |  | 143 | -575 |  | TWI011 | [7 empty] | TWI011
|  | TWI012 | 7.31 | 0.31 | 3.9% | 0.52 | 493 | -225 |  | TWI012 | 7.26 | 0.25 | 1.7% | 0.15 | 645 | -73 |  | TWI012 | 7.15 | 0.15 |  |  | 590 | -128 |  | TWI012
|  | TWI013 | 7.1 | 0.1 | 0% | 0 | 693 | -25 |  | TWI013 |  |  |  |  | 635 | -83 |  | TWI013 | [7 empty] | TWI013
|  | TWI014 | 7.05 | 0.05 | 0% | 0 | 670 | -48 |  | TWI014 |  |  |  |  | 779 | 61 |  | TWI014 | [7 empty] | TWI014
|  | TWI015 | 5.72 | -1.28 | 0% | 0 | 808 | 90 |  | TWI015 | 6.51 | -0.5 | 0% | 0 | 687 | -31 |  | TWI015 | 7.2 | 0.2 |  |  | 598 | -120 |  | TWI015
|  | TWI016 | 7.21 | 0.21 | 0% | 0 | 652 | -66 |  | TWI016 | 7.13 | 0.12 |  |  | 675 | -43 |  | TWI016 |  |  |  |  | 630 | -88 |  | TWI016
|  | TWI017 | 7.28 | 0.28 | 0% | 0 | 583 | -135 |  | TWI017 | 7.24 | 0.23 |  |  | 675 | -43 |  | TWI017 | 7.22 | 0.22 |  |  | 629 | -89 |  | TWI017 | 7.19 | 0.19 |  |  | 693 | -25
|  | TWI018 | blinking | #VALUE! | 0% | 0 | 571 | -147 |  | TWI018 | blinking | #VALUE! |  |  | 638 | -80 |  | TWI018 | blinking | #VALUE! |  |  | 613 | -105 |  | TWI018 | 7.2 | 0.2 |  |  | 541 | -177
|  | TWI019 | 7.37 | 0.37 | 0% | 0 | 547 | -171 |  | TWI019 | 7.17 | 0.16 |  |  | 612 | -106 |  | TWI019 |  |  |  |  | 605 | -113 |  | TWI019
|  | TWI020 | 7.13 | 0.13 | 0% | 0 | 570 | -148 |  | TWI020 |  |  |  |  | 681 | -37 |  | TWI020 | [7 empty] | TWI020
|  | TWI021 | 7.26 | 0.26 | 0% | 0 | 500 | -218 |  | TWI021 | 7.2 | 0.19 |  |  | 678 | -40 |  | TWI021 |  |  |  |  | 634 | -84 |  | TWI021
|  | TWI022 | 7.14 | 0.14 | 1.4% | 0.12 | 547 | -171 |  | TWI022 |  |  |  |  | 507 | -211 |  | TWI022 | [7 empty] | TWI022
|  | TWI023 | 7.05 | 0.05 | 0% | 0 | 583 | -135 |  | TWI023 |  |  |  |  | 379 | -339 |  | TWI023 | [7 empty] | TWI023
|  | TWI024 | [7 empty] | TWI024 | [7 empty] | TWI024 | [7 empty] | TWI024
|  | TWI025 | 7.13 | 0.13 | 0.5% | 0.04 | 744 | 26 |  | TWI025 |  |  |  |  | 771 | 53 |  | TWI025 | [7 empty] | TWI025
|  | NOTES: | TWI024 is not turning on and has been decomissioned. Rehydrated DO membranes and pH sensors where necessary. Calibrated all meters at 100% DO, 100 µS/cm, and 4.01 and 10.01 pH. Tested all meters at 0% DO, 718 µS/cm, and 7.01 pH. Errors with pH calibration of TWI015 and TWI018. Errors with DO calibration of TWI013. | [6 empty] | NOTES: | Replaced EC sensors, reset factory EC calibration, recalibrated and restested conductivity for TWI001, TWI002, TWI003, TWI004, TWI005, TWI006, TWI009, TWI010, TWI011, TWI013, TWI014, TWI020, TWI022, TWI023, and TWI025. Rehydrated DO membrane, recalibrated and retested DO for TWI007, TWI012, and TWI015. Rehydrated pH sensor, recalibrated and retested pH for TWI012 and TWI015. On 10/10/ 25, after soaking overnight, pH was recalibrated and retested for TWI008, TWI016, TWI017, TWI018, TWI019, and TWI021. Reset to factory calibration for conductivity, recalibrated and retested for TWI007, TWI008, TWI012, TWI015, TWI016, TWI017, TWI018, TWI019, and TWI021. | [6 empty] | NOTES: | Replaced EC sensors, reset to factory calibration, recalibrated and retested for TWI007, TWI008, TWI012, TWI015, TWI016, TWI017, TWI018, TWI019, and TWI021. Rehydrated DO membrane, reset to factory DO calibration, recalibrated and retested for TWI007. Reset to factory calibration, recalibrated and retested pH for TWI008, TWI012, TWI015, TWI017, and TWI018. | [6 empty] | NOTES: | Replaced DO membrane, reset to factory calibration, recalibrated and restested for TWI007. Failed again (5%, 0.43 ppm), so replaced DO sensor, installed DO membrane, reset to factory calibration, calibrated and retested TWI007 (round 5). Reset to factory conductivity calibration, recalibrated & retested for TWI007. Failed again (725), repeated recalibration and retesting for TWI007 (round 5). Replaced pH sensors on TWI017 and TWI018. Allowed to soak until Tuesday. On 10/14/25, reset to factory pH calibration. recalibrated and retested for TWI017 and TWI018. Recalibated and retested for conductivity on TWI017 & TWI018.
Note on AE112: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Tested 7/9
Note on AD138: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Round 5

## 2024 Testing
Meter Testing Procedures |  |  | JANUARY | [11 empty] | JANUARY | [8 empty] | JANUARY
| 1. After performing maintenance checks and replacing DO membranes, calibrate meters using volunteer methods |  | Round 1 | [11 empty] | Round 2 | [8 empty] | Round 3
|  |  | Test Date: | 2024-01-12 | [10 empty] | Test Date: | 2024-01-17 | [7 empty] | Test Date:
|  |  | Standards: | Water Temp | pH | DO% | DO ppm | Cond. |  |  | ± 1% cond. |  |  |  | Standards: | Water Temp | pH | DO% | DO ppm | Cond. |  |  |  | Standards: | Water Temp | pH | DO% | DO ppm | Cond. | ± 1% cond.
|  |  |  | 19.91 | 6.88 | 0 | 0 | 4523 |  |  | 45.23 | [5 empty] | 6.88 | 0 | 0 | [7 empty] | 0 | 0
|  |  | Meter | Water Temp | pH | pH diff. | DO% | DO ppm |  |  | Cond. |  | Cond. Diff. |  | Meter | Water Temp | pH | pH diff. | DO% | DO ppm |  | Cond. Diff. |  | Meter | Water Temp | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
| 2. Set up meter tests for DO, pH and conductivity. |  | TWI001 | 19.56 | 6.78 | -0.1 | 5.5% | 0.48 |  |  | 4347 |  | -176 |  | TWI001 |  |  |  | 12.5% | 1.08 |  |  |  | TWI001 |  |  |  | 12.5% | 1.08
|  |  | TWI002 | 19.83 | 7.15 | 0.27 | 0% | 0 |  |  | 4229 |  | -294 |  | TWI002 |  | 7.11 | 0.23 | [5 empty] | TWI002
|  | a. Set up two beakers with newly mixed 0% DO solution | TWI003 | 19.88 | 6.87 | -0.01 | 19.5% | 1.7 |  |  | 3170 |  | -1353 |  | TWI003 |  |  |  | 21.7% | 1.87 |  |  |  | TWI003 |  |  |  | 21.7% | 1.87
|  |  | TWI004 | [11 empty] | TWI004 | [8 empty] | TWI004
|  |  | TWI005 | 19.85 | 7.06 | 0.18 | 0% | 0 |  |  | 1852 |  | -2671 |  | TWI005 | [8 empty] | TWI005
|  | b. Set up two beakers with pH 7.01 buffer | TWI006 | 20.06 | 6.8 | -0.08 | 0% | 0 |  |  | 5119 |  | 596 |  | TWI006 | [8 empty] | TWI006
|  |  | TWI007 | 19.6 | 6.86 | -0.02 | 0% | 0 |  |  | 4711 |  | 188 |  | TWI007 | [8 empty] | TWI007
|  | c. Set up two beakers with conductivity solution ideally between 100 - ~700 μS/cm | TWI008 | 19.59 | 6.99 | 0.11 | 8.2% | 0.7 |  |  | 4977 |  | 454 |  | TWI008 |  |  |  | 0% | 0 |  |  |  | TWI008
|  |  | TWI009 | 19.85 | 6.98 | 0.1 | 0% | 0 |  |  | 3734 |  | -789 |  | TWI009 | [8 empty] | TWI009
|  |  | TWI010 | 19.82 | 6.78 | -0.1 | 0% | 0 |  |  | 4433 |  | -90 |  | TWI010 | [8 empty] | TWI010
| 3. Test meters against known DO, pH and conductivity standards. Write results in Quarterly Meter Testing sheet. |  | TWI011 | 20.33 | 6.87 | -0.01 | 0% | 0 |  |  | 908 |  | -3615 |  | TWI011 | [8 empty] | TWI011
|  |  | TWI012 | 19.62 | 6.87 | -0.01 | 0% | 0 |  |  | 4598 |  | 75 |  | TWI012 | [8 empty] | TWI012
|  |  | TWI013 | 19.79 | 6.93 | 0.05 | 0% | 0 |  |  | 4846 |  | 323 |  | TWI013 | [8 empty] | TWI013
|  |  | TWI014 |  | 6.93 | 0.05 | 1.3% | 0.12 |  |  | 3724 |  | -799 |  | TWI014 |  |  |  | 0% | 0 |  |  |  | TWI014
|  | a. Quarterly Meter Testing sheet | TWI015 |  | 7.06 | 0.18 | 0% | 0 |  |  | 5100 |  | 577 |  | TWI015 | [8 empty] | TWI015
|  |  | TWI016 |  | 6.8 | -0.08 | 0% | 0 |  |  | 4872 |  | 349 |  | TWI016 | [8 empty] | TWI016
|  |  | TWI017 |  | 6.98 | 0.1 | 0% | 0 |  |  | 4848 |  | 325 |  | TWI017 | [8 empty] | TWI017
|  |  | TWI018 | 19.78 | 6.56 | -0.32 | 0% | 0 |  |  | 4198 |  | -325 |  | TWI018 |  | 6.52 | -0.36 | [5 empty] | TWI018
|  |  | TWI019 | 19.88 | 6.86 | -0.02 | 0% | 0 |  |  | 4872 |  | 349 |  | TWI019 | [8 empty] | TWI019
|  |  | TWI020 | 19.67 | 6.89 | 0.01 | 0% | 0 |  |  | 2495 |  | -2028 |  | TWI020 | [8 empty] | TWI020
|  |  | TWI021 | 19.87 | 6.77 | -0.11 | 0% | 0 |  |  | 4641 |  | 118 |  | TWI021 | [8 empty] | TWI021
|  |  | TWI022 | 19.84 | 6.89 | 0.01 | 0% | 0 |  |  | 5068 |  | 545 |  | TWI022 | [8 empty] | TWI022
|  |  | TWI023 | 19.73 | 7.79 | 0.91 | 1% | 0.08 |  |  | 4495 |  | -28 |  | TWI023 |  | "check sensor" | NA | 0% | 0 |  |  |  | TWI023
|  |  | TWI024 | 19.87 | 6.66 | -0.22 | 1.8% | 0.14 |  |  | 4678 |  | 155 |  | TWI024 |  | 6.64 | -0.24 | 7.6% | 0.65 |  |  |  | TWI024 |  |  |  | 7.6% | 0.65
|  |  | TWI025 | 19.77 | 6.93 | 0.05 | 0% | 0 |  |  | 4595 |  | 72 |  | TWI025 | [8 empty] | TWI025
|  |  | NOTES: | TWI002, TWI018, TWI023 & TWI024 were recalibrated at 4 and 7 pH and then retested. TWI001, TWI003, TWI008, TWI014, TWI023, and TWI024 were all recalibrated at 100% and 0% DO then restested using a new aliquot of 0% DO solution. Test conductivity is far outside normal range of field measurments, so will ignore test results for this round and purchase more appropriate standard for next round of maintenance and testing. | [10 empty] | NOTES: | After recalibrating pH, TWI002, TWI018, TWI023 and TWI024 still failed the pH test. Sensors must be replaced for all four units. After recalibrating DO, TWI001, TWI003 and TWI024 still failed the DO test. DO membranes need to be replaced, then units need to be recalibrated and retested to determine if sensors need to be replaced. | [7 empty] | NOTES:
|  |  | APRIL | [13 empty] | APRIL | [8 empty] | APRIL | [8 empty] | APRIL
|  |  | Round 1 | [13 empty] | Round 2 | [8 empty] | Round 3 | [8 empty] | Round 4
|  |  | Test Date: | 2024-04-10 | [12 empty] | Test Date: | 2024-04-10 | [7 empty] | Test Date: | 2024-04-10 | [7 empty] | Test Date: | 2024-04-18
|  |  | Standards: | Water Temp | pH | DO% | DO ppm | [9 empty] | Standards: | Water Temp | pH | DO% | DO ppm |  |  |  |  | Standards: | Water Temp | pH | DO% | DO ppm |  |  |  |  | Standards: | Water Temp | pH | DO% | DO ppm
[5 empty] | 7.01 | 0 | 0 | [11 empty] | 7.01 | 0 | 0 | [6 empty] | 7.01 | 0 | 0 | [6 empty] | 7.01 | 0 | 0
|  |  | Meter | Water Temp | pH | pH diff. | DO% | DO ppm |  |  | YSI Cond |  | Cond | ± 1% cond. | Cond. Diff. |  | Meter | Water Temp | pH | pH diff. | DO% | DO ppm | Cond | Cond. Diff. |  | Meter | Water Temp | pH | pH diff. | DO% | DO ppm | Cond | Cond. Diff. |  | Meter | Water Temp | pH | pH diff. | DO% | DO ppm | Cond | Cond. Diff.
|  |  | TWI001 |  | 7.04 | 0.03 | 16.2% | 1.37 |  |  | 588 |  | 644 | 5.88 | 56 |  | TWI001 |  |  | -7.01 | 16.3% | 1.33 |  |  |  | TWI001 |  |  | -7.01 | 19.6% | 1.61 |  |  |  | TWI001 |  |  | -7.01 | 0.7% | 0.06
|  |  | TWI002 |  | 7.29 | 0.28 | 0% | 0 |  |  | 588 |  | 647 | 5.88 | 59 |  | TWI002 |  |  | -7.01 | [5 empty] | TWI002 |  | 7.19 | 0.18 | [5 empty] | TWI002 |  |  | -7.01
|  |  | TWI003 |  | 7.08 | 0.07 | 27% | 2.3 |  |  | 659 |  | 765 | 6.59 | 106 |  | TWI003 |  |  | -7.01 | 27.4% | 2.27 |  |  |  | TWI003 |  |  | -7.01 | 1.6% | 0.13 |  |  |  | TWI003 |  |  | -7.01 | 3.3% | 0.27
|  |  | TWI004 |  | 6.99 | -0.02 | 2% | 0.17 |  |  | 678 |  | 459 | 6.78 | 219 |  | TWI004 |  |  | -7.01 | 3.3% | 0.27 |  |  |  | TWI004 |  |  | -7.01 | [5 empty] | TWI004 |  |  | -7.01 | climbing
|  |  | TWI005 |  | "check sensor" | NA | 0% | 0 |  |  | 882 |  | 544 | 8.82 | 338 |  | TWI005 |  | 7.2 | 0.19 | [5 empty] | TWI005 |  |  | -7.01 | [5 empty] | TWI005 |  |  | -7.01
|  |  | TWI006 |  | 7.02 | 0.01 | 0% | 0 |  |  | 696 |  | 772 | 6.96 | 76 |  | TWI006 |  |  | -7.01 | [5 empty] | TWI006 |  |  | -7.01 | [5 empty] | TWI006 |  |  | -7.01
|  |  | TWI007 |  | "check sensor" | NA | 0% | 0 |  |  | 910 |  | 1533 | 9.1 | 623 |  | TWI007 |  |  | -7.01 | [5 empty] | TWI007 |  | 7.2 | 0.19 | [5 empty] | TWI007 |  |  | -7.01
|  |  | TWI008 |  | "check sensor" | NA | 0% | 0 |  |  | 924 |  | 1003 | 9.24 | 79 |  | TWI008 |  | 7.09 | 0.08 | [5 empty] | TWI008 |  |  | -7.01 | [5 empty] | TWI008 |  |  | -7.01
|  |  | TWI009 |  | 7.1 | 0.09 | 0% | 0 |  |  | 702 |  | 771 | 7.02 | 69 |  | TWI009 |  |  | -7.01 | [5 empty] | TWI009 |  |  | -7.01 | [5 empty] | TWI009 |  |  | -7.01
|  |  | TWI010 |  | 7.07 | 0.06 | 0% | 0 |  |  | 705 |  | 665 | 7.05 | 40 |  | TWI010 |  |  | -7.01 | [5 empty] | TWI010 |  |  | -7.01 | [5 empty] | TWI010 |  |  | -7.01
|  |  | TWI011 |  | 6.88 | -0.13 | 0% | 0 |  |  | 707 |  | 269 | 7.07 | 438 |  | TWI011 |  |  | -7.01 | [5 empty] | TWI011 |  |  | -7.01 | [5 empty] | TWI011 |  |  | -7.01
|  |  | TWI012 |  | 7.07 | 0.06 | 0% | 0 |  |  | 723 |  | 732 | 7.23 | 9 |  | TWI012 |  |  | -7.01 | [5 empty] | TWI012 |  |  | -7.01 | [5 empty] | TWI012 |  |  | -7.01
|  |  | TWI013 |  | "check sensor" | NA | 0% | 0 |  |  | 776 |  | 867 | 7.76 | 91 |  | TWI013 |  | 7.1 | 0.09 | [5 empty] | TWI013 |  |  | -7.01 | [5 empty] | TWI013 |  |  | -7.01
|  |  | TWI014 |  | 7.01 | 0 | 5.4% | 0.46 |  |  | 745 |  | 273 | 7.45 | 472 |  | TWI014 |  |  | -7.01 | 24% | 0.2 |  |  |  | TWI014 |  |  | -7.01 | [5 empty] | TWI014 |  |  | -7.01
|  |  | TWI015 |  | 7.21 | 0.2 | 0% | 0 |  |  | 734 |  | 857 | 7.34 | 123 |  | TWI015 |  |  | -7.01 | [5 empty] | TWI015 |  |  | -7.01 | [5 empty] | TWI015 |  |  | -7.01
|  |  | TWI016 |  |  | -7.01 | [7 empty] | 0 | 0 |  | TWI016 |  |  | -7.01 | [5 empty] | TWI016 |  | 7.26 | 0.25 | [5 empty] | TWI016 |  | 7.02 | 0.01
|  |  | TWI017 |  | 7.11 | 0.1 | 0% | 0 |  |  | 719 |  | 861 | 7.19 | 142 |  | TWI017 |  |  | -7.01 | [5 empty] | TWI017 |  |  | -7.01 | [5 empty] | TWI017 |  |  | -7.01
|  |  | TWI018 |  | 6.72 | -0.29 | 0% | 0 |  |  | 730 |  | 1029 | 7.3 | 299 |  | TWI018 |  |  | -7.01 | [5 empty] | TWI018 |  | 7.16 | 0.15 | [5 empty] | TWI018 |  |  | -7.01
|  |  | TWI019 |  | 7.21 | 0.2 | 0% | 0 |  |  | 772 |  | 802 | 7.72 | 30 |  | TWI019 |  |  | -7.01 | [5 empty] | TWI019 |  |  | -7.01 | [5 empty] | TWI019 |  |  | -7.01
|  |  | TWI020 |  |  | -7.01 | [7 empty] | 0 | 0 |  | TWI020 |  |  | -7.01 | [5 empty] | TWI020 |  |  | -7.01 | [5 empty] | TWI020 |  |  | -7.01
|  |  | TWI021 |  | "check sensor" | NA | 0% | 0 |  |  | 780 |  | 758 | 7.8 | 22 |  | TWI021 |  | 7.04 | 0.03 | [5 empty] | TWI021 |  |  | -7.01 | [5 empty] | TWI021 |  |  | -7.01
|  |  | TWI022 |  | 7.12 | 0.11 | 0% | 0 |  |  | 785 |  | 880 | 7.85 | 95 |  | TWI022 |  |  | -7.01 | [5 empty] | TWI022 |  |  | -7.01 | [5 empty] | TWI022 |  |  | -7.01
|  |  | TWI023 |  | "check sensor" | NA | 23.8% | 1.98 |  |  | 782 |  | 770 | 7.82 | 12 |  | TWI023 |  |  | -7.01 | 0% | 0 |  |  |  | TWI023 |  | 7.2 | 0.19 | [5 empty] | TWI023 |  |  | -7.01
|  |  | TWI024 |  | 6.83 | -0.18 | 100.8% | 8.93 |  |  | 815 |  | 862 | 8.15 | 47 |  | TWI024 |  |  | -7.01 | 2% | 0.16 |  |  |  | TWI024 |  |  | -7.01 | [5 empty] | TWI024 |  |  | -7.01 | climbing
|  |  | TWI025 |  | "check sensor" | NA | 0% | 0 |  |  | 784 |  | 904 | 7.84 | 120 |  | TWI025 |  | "check sensor" | NA | 0% | 0 |  |  |  | TWI025 |  | 7.19 | 0.18 | [5 empty] | TWI025 |  |  | -7.01
|  |  | NOTES: | TWI020 was not returned for this round of maintenance. The protective cover is stuck on TWI016. We were unable to perform maintenance or testing. All other meters were calibrated as volunteers are trained. TWI005, TWI007, TWI008, TWI013, TWI021, TWI023 and TWI025 showed "check sensor" errors on pH calibration. We used diluted quick cal solution to test conductivity and compared against simultaneous YSI readings. | [12 empty] | NOTES: | New DO membrane caps were installed and recalibrated on TWI001, TWI003, TWI004, TWI014, TWI023, and TWI024 prior to retesting. All but TWI023 failed second round testing for DO. New pH sensors were installed and recalibrated on TWI005, TWI007, TWI008, TWI013, TWI021, TWI023 and TWI025 prior to second round testing. All but TWI025 passed second round testing for pH. TWI025 was left to rehydrate pH sensor overnight before third round testing. Forgot to restest pH on TWI023. | [7 empty] | NOTES: | We only had two new DO sensors on hand, so DO sensors were replaced and recalibrated on TWI001 and TWI004 only. Both failed third round testing for DO. TWI001, TWI003, TWI004, TWI014, TWI023 and TWI024 have been decomissioned for now. pH needs to be tested on TWI023. The rest should be set to factory calibrations (per the manual), then recalibrated and restested. On 4/16 DO membrane cap was rehydrated on TWI014. It was reset to factory DO calibration and then recalibrated at 100% DO. We were unable to retest DO before it was returned to the volunteer on 4/17 | [7 empty] | NOTES: | Reset to factory pH calibration for TWI002, TWI007, TWI018 and TWI023. They were then recalibrated at 4.01 and 10.00 pH and tested at 7.01 pH. All but TWI016 passed. Reset to factory DO calibration for TWI001, TWI003, TWI004, TWI024. Tested with new 0% DO solution which came up at 0.07% on the YSI meter. All failed the DO test with only two settling and the other two climbing continuously.TWI101 had DO cap replaced and was then calibrated and tested. Unable to calibrate pH due to "contaminated buffer" error. TWI016 was left in 4.01 pH buffer for several hours and then recalibrated and retested. It passed.
|  |  | JULY | [10 empty] | JULY | [7 empty] | JULY | [7 empty] | JULY
|  |  | Round 1 | [10 empty] | Round 2 | [7 empty] | Round 3 | [7 empty] | Round 4
|  |  | Test Date: | 2024-08-06 | [9 empty] | Test Date: | 2024-08-06 | [6 empty] | Test Date: | 2024-08-07 | [6 empty] | Test Date: | 2024-08-08
|  |  | Standards: | pH | DO% | DO ppm | Cond. | ± 10% cond. | [5 empty] | Standards: | pH | DO% | DO ppm | Cond. | ± 10% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ± 10% cond. |  |  | Standards: | pH | DO% | DO ppm | Cond. | ± 10% cond.
|  |  |  | 7 | 0 | 0 | 74 | 7.4 | [6 empty] | 7 | 0 | 0 | 74 | 7.4 |  |  |  | 7 | 0 | 0 | 74 | 7.4 |  |  |  | 7 | 0 | 0 | 74 | 7.4
|  |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. |  |  | Cond. Diff. |  |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  |  | TWI001 | 7.1 | 0.1 | 0% |  | 70 |  |  | 4 |  |  | TWI001 |  | -7 |  |  |  | 7.4 |  | TWI001 |  | -7 |  |  |  | 74 |  | TWI001 |  | -7 |  |  |  | 7.4
|  |  | TWI002 | 7.18 | 0.18 | 3.3% |  | 72 |  |  | 2 |  |  | TWI002 |  | -7 |  |  |  | 7.4 |  | TWI002 | 7.12 | 0.12 | 0% |  | 74 | 0 |  | TWI002 |  | -7 |  |  |  | 7.4
|  |  | TWI003 | 7.06 | 0.06 | 1% |  | 82 |  |  | 8 |  |  | TWI003 |  | -7 | 0.8% |  |  | 7.4 |  | TWI003 | 7.02 | 0.02 | 0.7% |  | 79 | 5 |  | TWI003 |  | -7 |  |  |  | 7.4
|  |  | TWI004 | 6.82 | -0.18 | 26.7% |  | 62 |  |  | 12 |  |  | TWI004 |  | -7 |  |  |  | 7.4 |  | TWI004 | 7.2 | 0.2 | unstable |  | 67 | 7 |  | TWI004 |  | -7 | 1% |  |  | 7.4
|  |  | TWI005 | 7.19 | 0.19 | 0% |  | 74 |  |  | 0 |  |  | TWI005 |  | -7 |  |  |  | 7.4 |  | TWI005 |  | -7 |  |  |  | 74 |  | TWI005 |  | -7 |  |  |  | 7.4
|  |  | TWI006 | 7.15 | 0.15 | 0% |  | 79 |  |  | 5 |  |  | TWI006 |  | -7 |  |  |  | 7.4 |  | TWI006 |  | -7 |  |  |  | 74 |  | TWI006 |  | -7 |  |  |  | 7.4
|  |  | TWI007 | 6.56 | -0.44 | 14% |  | 95 |  |  | 21 |  |  | TWI007 |  | -7 | 5.7% |  |  | 7.4 |  | TWI007 | 6.43 | -0.57 | 1.3% |  | 69 | 5 |  | TWI007 | 7.2 | 0.2 | 1% |  |  | 7.4
|  |  | TWI008 | 7.14 | 0.14 | 0% |  | 75 |  |  | 1 |  |  | TWI008 |  | -7 |  |  |  | 7.4 |  | TWI008 |  | -7 |  |  |  | 74 |  | TWI008 |  | -7 |  |  |  | 7.4
|  |  | TWI009 | 7.16 | 0.16 | 0% |  | 73 |  |  | 1 |  |  | TWI009 |  | -7 |  |  |  | 7.4 |  | TWI009 |  | -7 |  |  |  | 74 |  | TWI009 |  | -7 |  |  |  | 7.4
|  |  | TWI010 | 7.13 | 0.13 | 0% |  | 71 |  |  | 3 |  |  | TWI010 |  | -7 |  |  |  | 7.4 |  | TWI010 |  | -7 |  |  |  | 74 |  | TWI010 |  | -7 |  |  |  | 7.4
|  |  | TWI011 | 7.36 | 0.36 | 0% |  | 80 |  |  | 6 |  |  | TWI011 |  | -7 |  |  |  | 7.4 |  | TWI011 | 7.33 | 0.33 | 0% |  | 79 | 5 |  | TWI011 | 7.19 | 0.19 |  |  |  | 7.4
|  |  | TWI012 | 6.95 | -0.05 | 1.5% |  | 105 |  |  | 31 |  |  | TWI012 |  | -7 | 0.7% |  |  | 7.4 |  | TWI012 | 7.03 | 0.03 | 1% |  | 70 | 4 |  | TWI012 |  | -7 |  |  |  | 7.4
|  |  | TWI013 | 7.17 | 0.17 | 0% |  | 110 |  |  | 36 |  |  | TWI013 |  | -7 |  |  |  | 7.4 |  | TWI013 | 7.12 | 0.12 | 0% |  | 75 | 1 |  | TWI013 |  | -7 |  |  |  | 7.4
|  |  | TWI014 | 7.22 | 0.22 | 4.5% |  | 77 |  |  | 3 |  |  | TWI014 |  | -7 | 4% |  |  | 7.4 |  | TWI014 | 7.15 | 0.15 | 3.9% |  | 67 | 7 |  | TWI014 |  | -7 | 0% |  |  | 7.4
|  |  | TWI015 | 7.02 | 0.02 | 0% |  | 75 |  |  | 1 |  |  | TWI015 |  | -7 |  |  |  | 7.4 |  | TWI015 |  | -7 |  |  |  | 74 |  | TWI015 |  | -7 |  |  |  | 7.4
|  |  | TWI016 | 7.16 | 0.16 | 0% |  | 71 |  |  | 3 |  |  | TWI016 |  | -7 |  |  |  | 7.4 |  | TWI016 |  | -7 |  |  |  | 74 |  | TWI016 |  | -7 |  |  |  | 7.4
|  |  | TWI017 | 7.15 | 0.15 | 0% |  | 81 |  |  | 7 |  |  | TWI017 |  | -7 |  |  |  | 7.4 |  | TWI017 |  | -7 |  |  |  | 74 |  | TWI017 |  | -7 |  |  |  | 7.4
|  |  | TWI018 | 7.18 | 0.18 | 0% |  | 79 |  |  | 5 |  |  | TWI018 |  | -7 |  |  |  | 7.4 |  | TWI018 |  | -7 |  |  |  | 74 |  | TWI018 |  | -7 |  |  |  | 7.4
|  |  | TWI019 | 6.86 | -0.14 | 0% |  | 76 |  |  | 2 |  |  | TWI019 |  | -7 |  |  |  | 7.4 |  | TWI019 |  | -7 |  |  |  | 74 |  | TWI019 |  | -7 |  |  |  | 7.4
|  |  | TWI020 | 7.17 | 0.17 | 0% |  | 67 |  |  | 7 |  |  | TWI020 |  | -7 |  |  |  | 7.4 |  | TWI020 |  | -7 |  |  |  | 74 |  | TWI020 |  | -7 |  |  |  | 7.4
|  |  | TWI021 | 7.08 | 0.08 | 0% |  | 77 |  |  | 3 |  |  | TWI021 |  | -7 |  |  |  | 7.4 |  | TWI021 |  | -7 |  |  |  | 74 |  | TWI021 |  | -7 |  |  |  | 7.4
|  |  | TWI022 | 7.12 | 0.12 | 0% |  | 83 |  |  | 9 |  |  | TWI022 |  | -7 |  |  |  | 7.4 |  | TWI022 | 7.08 | 0.08 | 0% |  | 80 | 6 |  | TWI022 |  | -7 |  |  |  | 7.4
|  |  | TWI023 | 7.11 | 0.11 | 0% |  | 75 |  |  | 1 |  |  | TWI023 |  | -7 |  |  |  | 7.4 |  | TWI023 |  | -7 |  |  |  | 74 |  | TWI023 |  | -7 |  |  |  | 7.4
|  |  | TWI024 | 6.89 | -0.11 | 0% |  | 87 |  |  | 13 |  |  | TWI024 |  | -7 |  |  |  | 7.4 |  | TWI024 | 6.89 | -0.11 | 0% |  | 71 | 3 |  | TWI024 |  | -7 |  |  |  | 7.4
|  |  | TWI025 | 6.99 | -0.01 | 17.1% |  | 77 |  |  | 3 |  |  | TWI025 |  | -7 | 337.4% |  |  | 7.4 |  | TWI025 | 7.02 | 0.02 | unstable |  | 67 | 7 |  | TWI025 |  | -7 | 0% |  |  | 7.4
|  |  | NOTES: | *New test tracking form includes calibration log.* Replaced all DO caps. All meters were calibrated as volunteers are trained for DO and EC. All meters were calibrated at 4 and 10 pH. TWI004 and TWI019 showed "wrong standard" errors on pH calibration but passed the pH test at 7 pH. The three large meters were also calibrated and tested. Calibration for DO, pH and EC are the same for the big meters. For turbidity, they were calibrated at 0 (using DI water) and 20 NTU and then tested at 100 NTUs. The pH sensor on TWI101 was replaced before calibration, and the meter could not detect the new pH sensor (possibly due to liquid or corrosion in the connection port). | [9 empty] | NOTES: | Changed acceptible error for conductivity to 10%. Replaced DO sensor for TWI007 and TWI025. Rehydrated DO membrane cap for TWI003, TWI012, and TWI014. | [6 empty] | NOTES: | Allowed TWI007, TWI011, TWI014 and TWI025 to soak in pH storage solution overnighted. Changed acceptible error for DO to 1%. Changed EC sensors for TWI007, TWI012, TWI013 & reset to facrory calibration before retesting. Changed DO sensor for TWI004 and reset to factory calibration before retesting. | [6 empty] | NOTES: | On 8/7: Reset to factory calibration for DO & restested at 0% for TWI007. Passed. On 8/8: Recalibrated at 10 and 4 and tested at 7 pH for TWI007 and TWI011. Changed DO membrane caps for TWI004, TWI014, and TWI025. Reset to facory calibration for DO on TWI004 and TWI025. Calibrated DO at 100% and 0% for TWI014 and TWI025.
|  |  | OCTOBER | [10 empty] | OCTOBER
|  |  | Round 1 | [10 empty] | Round 2
|  |  | Test Date: | 2024-10-10 | [9 empty] | Test Date: | 2024-10-10
|  |  | Standards: | pH | DO% | DO ppm | Cond. | ± 10% cond. | Cond. | ± 10% cond. |  |  |  | Standards: | pH | DO% | DO ppm | Cond. | ± 10% cond.
|  |  |  | 7 | 0 | 0 | 74 | 7.4 | 447 | 67.05 |  |  |  |  | 7 | 0 | 0 | 74 | 7.4
|  |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff. | Cond. | Cond. Diff |  |  | Meter | pH | pH diff. | DO% | DO ppm | Cond. | Cond. Diff.
|  |  | TWI001 | 6.83 | -0.17 | 2.3% | 0.19 | 70 | 4 | 407 | 40 | 5.714% | 9.828% | TWI001
|  |  | TWI002 | 7.02 | 0.02 | 0% | 0 | 75 | 1 | 395 | 52 | 1.333% | 13.16% | TWI002
|  |  | TWI003 | 7.03 | 0.03 | 0.9% | 0.06 | 73 | 1 | 472 | 25 | 1.37% | 5.297% | TWI003
|  |  | TWI004 | 6.92 | -0.08 | 2.3% | 0.19 | 80 | 6 | 429 | 18 | 7.5% | 4.196% | TWI004
|  |  | TWI005 | 7.27 | 0.27 | 0% | 0 | 72 | 2 | 388 | 59 | 2.778% | 15.21% | TWI005 | 7.27 | 0.27
|  |  | TWI006 | 6.95 | -0.05 | 0% | 0 | 81 | 7 | 471 | 24 | 8.642% | 5.096% | TWI006
|  |  | TWI007 | 7 | 0 | 4% | 0.34 | 69 | 5 | 367 | 80 | 7.246% | 21.8% | TWI007
|  |  | TWI008 | 7.11 | 0.11 | 0% | 0 | 72 | 2 | 434 | 13 | 2.778% | 2.995% | TWI008
|  |  | TWI009 | 7.05 | 0.05 | 0% | 0 | 77 | 3 | 438 | 9 | 3.896% | 2.055% | TWI009
|  |  | TWI010 | 7.24 | 0.24 | 0% | 0 | 72 | 2 | 416 | 31 | 2.778% | 7.452% | TWI010 | 7.25 | 0.25
|  |  | TWI011 | 7.11 | 0.11 | 0% | 0 | 81 | 7 | 365 | 82 | 8.642% | 22.47% | TWI011
|  |  | TWI012 | 7.45 | 0.45 | 2% | 0.17 | 73 | 1 | 414 | 33 | 1.37% | 7.971% | TWI012 | 7.36 | 0.36
|  |  | TWI013 | 7.18 | 0.18 | 0% | 0 | 67 | 7 | 476 | 29 | 10.45% | 6.092% | TWI013
|  |  | TWI014 | 7.2 | 0.2 | 0% | 0 | 62 | 12 | 370 | 77 | 19.35% | 20.81% | TWI014
|  |  | TWI015 | 7.5 | 0.5 | 0% | 0 | 80 | 6 | 458 | 11 | 7.5% | 2.402% | TWI015 | 7.22 | 0.22
|  |  | TWI016 | 6.83 | -0.17 | 0% | 0 | 82 | 8 | 402 | 45 | 9.756% | 11.19% | TWI016
|  |  | TWI017 | 7.03 | 0.03 | 0% | 0 | 72 | 2 | 418 | 29 | 2.778% | 6.938% | TWI017
|  |  | TWI018 | 7.16 | 0.16 | 0% | 0 | 81 | 7 | 446 | 1 | 8.642% | 0.2242% | TWI018
|  |  | TWI019 | 7.18 | 0.18 | 0% | 0 | 70 | 4 | 430 | 17 | 5.714% | 3.953% | TWI019
|  |  | TWI020 | 7.13 | 0.13 | 0% | 0 | 65 | 9 | 377 | 70 | 13.85% | 18.57% | TWI020
|  |  | TWI021 | 7.06 | 0.06 | 0% | 0 | 74 | 0 | 421 | 26 | 0% | 6.176% | TWI021
|  |  | TWI022 | 7.11 | 0.11 | 0% | 0 | 72 | 2 | 431 | 16 | 2.778% | 3.712% | TWI022
|  |  | TWI023 | 7.19 | 0.19 | 0% | 0 | 75 | 1 | 412 | 35 | 1.333% | 8.495% | TWI023
|  |  | TWI024 | 7.15 | 0.15 | 0% | 0 | 75 | 1 | 420 | 27 | 1.333% | 6.429% | TWI024
|  |  | TWI025 | 7.16 | 0.16 | 0% | 0 | 76 | 2 | 378 | 69 | 2.632% | 18.25% | TWI025
|  |  | NOTES: | Replaced DO caps. All meters calibrated at 100% DO, $.01 & 10 pH & 100 µS/cm then tested at 7 pH, 0% DO & 74 & 447 µS/cm. TWI025 showed "wrong standard" error when calibrating DO but passed the DO test at 0%. TWI005, TWI010, TWI012 & TWI 015 failed the pH test. TWI001, TWI004, TWI007 & TWI012 failed the DO test. TWI007, TWI011, TWI014, TWI016, TWI020 & TWI025 failed the Conductivity test. | [9 empty] | NOTES: | TWI005, TWI010, TWI012 & TWI015 were reclaibrated at 4.01 & 1- pH and then retested at 7 pH. All failed but not be enough to change the sensors. *No further notes on testing or sensor replacement.* METER MAINTENANCE & TESTING SOP WAS WRITTEN AFTER THIS ROUND OF TESTING
Note on AC50: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Tested 4/18/24
Note on AC55: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Tested 4/18/24
Note on AC64: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Tested 4/18/24
Note on AC66: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Tested 4/18/24
Note on AC71: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Tested 4/18/24
Note on AC73: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Tested 4/18/24
Note on AJ96: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Tested 8/7/24